Traces:

| Trace | Start | Step | Stop | IF Bandwidth | Detector | Dwell time | Attenuation | Transducer | LISN | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Scan 1 | 30.0 MHz | 50.0 kHz | 300.0 MHz | 120kHz | QP | 1.0 s | Auto | TDF1, TDF2, TDF3 | None- |  |

Transducers:

TDF1: cdne20db

TDF2: Cable MC 38377 - GAUSS Instruments (DC - 3GHz)

TDF3: Testec BNC cable

Limit lines:

CISPR15ED9: CISPR 15 ED9 QP,CDNE, 30MHz-300MHz

Scan1: 30.0 MHz, 50.0 kHz, 300.0 MHz; IF:120kHz, 1.0 s QP, Att AutodB

| f | Mag [dBµV] | Limit | Diff | Name | Comment |
| --- | --- | --- | --- | --- | --- |
| 79.888 MHz | 45.43 | 55.87 | 10.43 | CISPR15ED9 |  |
| 137.494 MHz | 48.33 | 54.00 | 5.67 | CISPR15ED9 |  |
| 152.220 MHz | 54.31 | 54.00 | -0.31 | CISPR15ED9 | Failed |
| 153.086 MHz | 54.59 | 54.00 | -0.59 | CISPR15ED9 | Failed |
| 154.031 MHz | 54.71 | 54.00 | -0.71 | CISPR15ED9 | Failed |
| 154.583 MHz | 54.80 | 54.00 | -0.80 | CISPR15ED9 | Failed |
| 155.646 MHz | 54.82 | 54.00 | -0.82 | CISPR15ED9 | Failed |
| 155.882 MHz | 54.71 | 54.00 | -0.71 | CISPR15ED9 | Failed |
| 156.709 MHz | 54.57 | 54.00 | -0.57 | CISPR15ED9 | Failed |
| 157.732 MHz | 54.33 | 54.00 | -0.33 | CISPR15ED9 | Failed |
| 159.938 MHz | 54.28 | 54.00 | -0.28 | CISPR15ED9 | Failed |
| 160.292 MHz | 54.32 | 54.00 | -0.32 | CISPR15ED9 | Failed |
| 194.312 MHz | 38.07 | 54.00 | 15.93 | CISPR15ED9 |  |
| 266.723 MHz | 34.01 | 51.87 | 17.86 | CISPR15ED9 |  |