Traces:

| Trace | Start | Step | Stop | IF Bandwidth | Detector | Dwell time | Attenuation | Transducer | LISN | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Scan 1 | 30.0 MHz | 50.0 kHz | 300.0 MHz | 120kHz | QP | 1.0 s | Auto | TDF1, TDF2, TDF3 | None- |  |

Transducers:

TDF1: cdne20db

TDF2: Cable MC 38377 - GAUSS Instruments (DC - 3GHz)

TDF3: Testec BNC cable

Limit lines:

CISPR15ED9: CISPR 15 ED9 QP,CDNE, 30MHz-300MHz

Scan1: 30.0 MHz, 50.0 kHz, 300.0 MHz; IF:120kHz, 1.0 s QP, Att AutodB

| f | Mag [dBµV] | Limit | Diff | Name | Comment |
| --- | --- | --- | --- | --- | --- |
| 79.849 MHz | 45.99 | 55.87 | 9.88 | CISPR15ED9 |  |
| 137.809 MHz | 49.16 | 54.00 | 4.84 | CISPR15ED9 |  |
| 150.527 MHz | 54.16 | 54.00 | -0.16 | CISPR15ED9 | Failed |
| 151.196 MHz | 54.34 | 54.00 | -0.34 | CISPR15ED9 | Failed |
| 151.748 MHz | 54.37 | 54.00 | -0.37 | CISPR15ED9 | Failed |
| 152.535 MHz | 54.44 | 54.00 | -0.44 | CISPR15ED9 | Failed |
| 152.929 MHz | 54.55 | 54.00 | -0.55 | CISPR15ED9 | Failed |
| 153.401 MHz | 54.46 | 54.00 | -0.46 | CISPR15ED9 | Failed |
| 154.031 MHz | 54.41 | 54.00 | -0.41 | CISPR15ED9 | Failed |
| 154.858 MHz | 54.37 | 54.00 | -0.37 | CISPR15ED9 | Failed |
| 155.803 MHz | 54.35 | 54.00 | -0.35 | CISPR15ED9 | Failed |
| 156.197 MHz | 54.20 | 54.00 | -0.20 | CISPR15ED9 | Failed |
| 194.509 MHz | 39.74 | 54.00 | 14.26 | CISPR15ED9 |  |
| 266.447 MHz | 35.24 | 51.88 | 16.64 | CISPR15ED9 |  |